# VSX3000 Resolution, Area-Fidelity, Detectability, and Noise Characterization Procedure

Oct 4, 2026 · @G Neil

## 1. Purpose and scope

This procedure collects the data needed to measure four properties of the VSX3000 active binocular stereo sensor over 500–1000 mm. The sensor stays fixed. A robot carries the target, so each target pose is commanded, repeatable, and logged.

The four measurements are not independent. Procedure A supplies the noise level that sets the detection thresholds in B, C, and D. Procedure B supplies the edge spread function that predicts the area bias measured in C. Part I (acquisition) runs in the order A, B, C, D. Part II (analysis) runs in the same order, then E. E reuses data from B and C and needs no separate captures.

Procedure order and data dependencies · 5 acquisition series, 5 analyses

Shaded boxes are captures and open boxes are analyses. Registration poses feed every analysis, and the quick-look on C data sets the D trial levels (dashed).

Coordinate convention. Z runs along the left IR camera's optical axis, with its origin at that camera's optical center. H runs along image columns and V along image rows. H is assumed parallel to the stereo baseline; Step 4.5 checks this. The depth image is assumed to be registered to the left IR image; Step 4.5 checks this as well.

| ID | Measurand | Definition used in this procedure | Primary outputs |
| --- | --- | --- | --- |
| A | Noise vs Z | Temporal, fixed-pattern, and total depth deviation on a flat matte plate. Also fill rate, spatial correlation length, and depth quantization step. | σ_t(Z), σ_fp(Z), σ_tot(Z), bias(Z), fitted disparity noise σ_d, quantum q, correlation length |
| B-HV | Effective resolution in H and V | 10–90% rise distance and MTF50 of the depth edge spread function (ESF), for both edge polarities | Rise distance in mm and px for each Z, orientation, and polarity |
| B-Z | Effective resolution in Z | Smallest commanded Z step detected with 50% probability, for 1-px, 5×5 px, and 20×20 px patches. Also the staircase quantum and step-response gain. | δ_50(Z, patch), δZ_q(Z), gain |
| C | True vs sensed area | Sensed area at the half-height iso-contour compared with true area and with geometrically visible area, for disks and for cutouts | A_sensed/A_true vs D/p(Z); edge bias b(Z) |
| D | Minimum detectable size | Diameter and subtended angle at 50%, 10%, and 0% detection probability, corrected for false alarms, for disks and for cutouts | D_50, D_10, D_0 and θ_50, θ_10, θ_0 for each Z, with confidence intervals |
| E | Boundary detection bias | Near a true edge, how often the sensor reports a value versus a no-read, and whether it favors the near or the far surface | Read/no-read bias index, near/far preference, and outcome-probability profiles vs distance to the edge |

The outputs map onto the Tier-A forward-noise simulator parameters: σ_d maps to σ_j, the quantum to q, the measured k = f·B to k, and the correlation length to the IDW neighborhood N. Out of scope: surface-material dependence, ambient sunlight, multipath, and motion blur. Each of these can be added later as a repeat of A–D with one variable changed.

## 2. Parameters

Every arbitrary constant in this procedure is a named parameter listed here. The suggested values are starting points for the VSX3000 at 500–1000 mm. Values marked † depend on VSX3000 datasheet or SDK values I do not have, so they must be confirmed or replaced in Step 4.5.

| Name | Suggested value | Units | Used in | Meaning |
| --- | --- | --- | --- | --- |
| Sensor geometry |  |  |  |  |
| SENSOR_FX_PX, SENSOR_FY_PX | from SDK † | px | all | Left IR focal lengths |
| SENSOR_CX_PX, SENSOR_CY_PX | from SDK † | px | all | Left IR principal point |
| SENSOR_BASELINE_MM | from datasheet † | mm | C, E | Stereo baseline B |
| PROJECTOR_OFFSET_MM | from datasheet † | mm (H, V, Z) | C, E | Projector center relative to the left camera |
| DEPTH_LSB_MM | from SDK † | mm | A | Least significant bit of the depth output |
| Working volume |  |  |  |  |
| Z_MIN_MM / Z_MAX_MM | 500 / 1000 | mm | all | Range limits |
| Z_NOISE_STEP_MM | 50 | mm | A | Spacing of noise stations, giving 11 stations |
| Z_SHAPE_STATIONS_MM | 500, 625, 750, 875, 1000 | mm | B, C, D | Stations for edge, area, and detection tests |
| Z_REDUCED_STATIONS_MM | 500, 750, 1000 | mm | B-Z, D (0% series) | Subset used for the slowest tests |
| FIELD_OFFSET_FRACTION | 0.6 | of half-FOV | A, C | Off-axis field positions at (±f·W/2, ±f·H/2) from center |
| TILT_ANGLES_DEG | 0, 15, 30, 45 | deg | A | Plate tilts about H and about V |
| Capture control |  |  |  |  |
| WARMUP_MIN_MINUTES | 45 | min | 4 | Minimum powered time before any capture |
| WARMUP_DRIFT_WINDOW_MIN | 10 | min | 4 | Window over which drift is judged |
| WARMUP_DRIFT_FRACTION_OF_SIGMA | 0.1 | — | 4 | Allowed drift of the mean plane Z, as a fraction of σ_t at that Z |
| DRIFT_SENTINEL_INTERVAL_MIN | 60 | min | all | Interval between sentinel captures |
| ROBOT_SETTLE_TIME_S | 2.0 | s | all | Wait after motion before capture; Step 4.4 verifies it |
| FRAMES_PER_NOISE_STATION | 100 | frames | A | Frames per Z, field, and tilt pose |
| FRAMES_PER_EDGE_POSE | 30 | frames | B-HV | Frames per edge pose |
| FRAMES_PER_ZSTEP_POSE | 10 | frames | B-Z | Frames per Z-step pose |
| FRAMES_PER_AREA_POSE | 10 | frames | C | Frames per area pose |
| FRAMES_PER_DETECTION_TRIAL | 1 | frames | D | One frame per trial, giving single-frame detectability |
| Phase randomization |  |  |  |  |
| PHASE_JITTER_SPAN_PX | 8 | px at the station Z | B, C, D | Square span of random lateral offsets, larger than the projector dot pitch and the noise correlation length |
| PHASE_JITTER_POSES_EDGE | 25 | poses | B-HV | Offsets per edge station |
| PHASE_JITTER_POSES_AREA | 30 | poses | C | Offsets per area station |
| Z-step test |  |  |  |  |
| Z_STEP_LADDER_MM | 0.1, 0.2, 0.5, 1, 2, 4 | mm | B-Z | Commanded step sizes |
| Z_STEP_REPEATS | 10 | ABAB cycles | B-Z | Alternations for each step size |
| Z_STAIRCASE_SUBDIVISION | 10 | per quantum | B-Z | Fine-sweep points per expected quantum |
| ROBOT_REPEATABILITY_MM | 0.05 | mm | B-Z | Position repeatability of the robot (ISO 9283, the Section 3.1 requirement). The read-back pose is the step truth of B-Z, so a rung smaller than this is captured but flagged: its truth is no better than the robot itself |
| Z_STEP_APPROACH_OVERSHOOT_MM | 2 | mm | B-Z | Every Z visit is approached from below by this overshoot and a final move up onto the pose, so that backlash and compliance do not enter the A→B difference |
| Detection (D, also used in B-Z) |  |  |  |  |
| DETECTION_FALSE_ALARM_TARGET | 0.01 | per window | B-Z, D | Sets threshold τ from the blank-site distribution |
| DETECTION_MIN_CONNECTED_PX | 2 | px | D | Minimum connected region counted as a detection |
| DETECTION_WINDOW_MARGIN_PX | 3 | px | D | Search window radius equals D/2 in px plus this margin |
| DETECTION_LEVELS | 9 | diameters | D | Diameter levels per curve |
| DETECTION_LEVEL_LOW_FACTOR / HIGH_FACTOR | 0.3 / 2.5 | × pilot D_50 | D | Range spanned by the levels, wide enough to include the 0% and 10% points |
| DETECTION_TRIALS_PER_LEVEL | 60 | trials | D | Independent trials per level |
| DETECTION_ZERO_TRIALS | 300 | trials | D | Trials at the levels next to the 0% point |
| DETECTION_ZERO_PROBABILITY_BOUND | 0.01 | — | D | Upper bound that defines practically zero detection |
| CONFIDENCE_LEVEL | 0.95 | — | B, C, D, E | For all confidence intervals and bounds |
| BOOTSTRAP_RESAMPLES | 2000 | — | C, D, E | Resamples for confidence intervals |
| Boundary bias (E) |  |  |  |  |
| BOUNDARY_BAND_HALF_WIDTH_PX | 8 | px | E | Analysis band on each side of a true edge |
| BOUNDARY_BIN_WIDTH_PX | 0.25 | px | B, E | Signed-distance bin width, giving 4× oversampling |
| SURFACE_ASSIGNMENT_SIGMA_MULTIPLE | 3 | × σ_tot(Z) | E | A read within this distance of a reference plane is assigned to that plane |
| Targets |  |  |  |  |
| NOISE_PLATE_SIZE_MM | 400 × 400 | mm | A, B-Z | Uniform matte plate |
| PLATE_FLATNESS_MM | 0.05 | mm | all | Must be at most 0.25 × the smallest expected σ_tot |
| GAP_SMALL_MM / GAP_LARGE_MM | 15 / 60 | mm | B, C, D, E | Distance from front plate to back plate |
| EDGE_SLANT_DEG | 5 | deg | B-HV | Edge rotation from the image axes |
| EDGE_SQUARE_SIZE_MM | 160 | mm | B-HV | Side of the raised square and of the square window |
| CHAMFER_MARGIN_DEG | 10 | deg | 3 | Margin added to the worst-case ray angle |
| EDGE_LAND_MAX_MM | 0.2 | mm | 3 | Maximum residual flat land at a knife edge |
| DIAMETER_LADDER_RATIO | √2 | — | C, D | Ratio between successive diameters |
| DIAMETER_MIN_FOOTPRINT_FRACTION | 0.3 | × p(Z_MIN) | C, D | Smallest diameter, as a fraction of the pixel footprint at Z_MIN |
| DIAMETER_MAX_FOOTPRINT_MULTIPLE | 30 | × p(Z_MAX) | C | Largest diameter, as a multiple of the pixel footprint at Z_MAX |
| FEATURE_ISOLATION_PX | 30 | px at Z_MIN | C, D | Minimum edge-to-edge spacing between features |
| POST_DIAMETER_FRACTION_OF_D0 | 0.5 | × pilot D_0 | 3 | Disk support posts must be thinner than this |
| Registration |  |  |  |  |
| REGISTRATION_POSES | 30 | poses | 4 | Hand-eye poses spanning the volume |
| REGISTRATION_TILT_RANGE_DEG | ±20 | deg | 4 | Tilt range about H and V |
| REGISTRATION_FRAMES | 10 | frames | 4 | Frames per registration pose |
| REGISTRATION_RESIDUAL_ACCEPT_MM | 0.15 | mm RMS | 4 | Acceptance limit for the registration residual |

These relations are used throughout. p(Z) is the pixel footprint, θ the subtended angle, d the disparity, and δZ_q the depth quantum implied by a disparity quantum q:

p(Z) = \frac{Z}{f_x}, \qquad \theta = \frac{D}{Z}, \qquad D_{\mathrm{px}} = \frac{D\, f_x}{Z}, \qquad d = \frac{k}{Z},\; k = f_x B, \qquad \delta Z_q = \frac{q\, Z^{2}}{k}

## 3. Equipment and targets

Every target uses a two-plane construction: a front surface with knife edges, standing a known gap G in front of a back plate of the same finish. One geometry therefore serves the edge, area, detection, and boundary-bias tests. Fabrication tolerances are left to the fabricator. Instead, every feature is measured as built and logged.

### 3.1 Equipment

| Item | Requirement | Why |
| --- | --- | --- |
| VSX3000 sensor on a rigid stand | Stand mechanically separate from the robot base, on an isolated floor or table | Prevents robot motion from moving the sensor |
| 6-axis robot | Repeatability ≤ 0.05 mm (ISO 9283); payload above the heaviest target plus adapter | Commanded, repeatable target poses |
| Quick-change target adapter | Dowel-pinned, re-mount repeatability ≤ 0.02 mm | Targets can be swapped without re-registering |
| Enclosure or blackout curtains, IR light meter | Ambient IR held constant and logged | Ambient IR changes the noise level |
| Temperature loggers (sensor housing and air) | 1 sample/min | For drift attribution |
| Capture PC | VSX3000 SDK, LRVisionLibs MatCloud reader, robot interface that logs the actual pose read back from the robot | Synchronized capture and logging |

### 3.2 Targets

| ID | Target | Construction | Used in |
| --- | --- | --- | --- |
| T1 | Registration plate | Flat, about 400 × 300 mm, with a ChArUco or circle-grid pattern that is visible in the left IR image. Pattern geometry is known to ≤ 0.02 mm. | 4 |
| T2 | Noise plate | Uniform matte, NOISE_PLATE_SIZE_MM, flatness PLATE_FLATNESS_MM | A, B-Z, sentinels |
| T3a | Raised square | Square of side EDGE_SQUARE_SIZE_MM with knife edges, on hidden posts at G above a back plate. Spacers set G to GAP_SMALL_MM or GAP_LARGE_MM. | B-HV, E |
| T3b | Square window | Front plate with a square window of the same size and knife edges; back plate at G behind it | B-HV, E |
| T4-S / T4-L | Disk arrays, small and large | Back-beveled disks on thin posts at G above a back plate. The diameter ladder is split across two plates by size. Includes blank sites and post-only control sites. | C, D, E |
| T5-S / T5-L | Cutout arrays, small and large | Back-beveled holes in a front plate, with the back plate at G. Includes blank sites. | C, D, E |

Diameter ladder. Diameters run from DIAMETER_MIN_FOOTPRINT_FRACTION × p(Z_MIN) up to DIAMETER_MAX_FOOTPRINT_MULTIPLE × p(Z_MAX), in steps of DIAMETER_LADDER_RATIO. For example, at f_x ≈ 500 px (unconfirmed) this gives about 0.3 mm to 60 mm in about 16 steps. Features on a plate are spaced at least FEATURE_ISOLATION_PX apart at Z_MIN, so the sensor's spatial interpolation cannot couple neighboring features.

Blank and control sites. Each array has blank sites matching each search-window size. These give the false-alarm rate in D. Disk arrays also have post-only sites (a post with no disk). These show whether the support post itself is detected.

Surface finish. All front and back surfaces use one finish, for example bead-blasted aluminum or a matte coating, with mid-range IR reflectance. Record the finish and, if possible, its reflectance at the projector wavelength. A reflectance difference between plates would bias the edge and area results.

### 3.3 Chamfered (knife-edge) boundaries

Every boundary that defines an edge, disk, or cutout must be chamfered from the back. Then no camera or projector ray can strike the boundary's side wall, and the sensor sees only the front face and the back plate. An unchamfered wall would be seen as a third surface by one camera and not the other, which corrupts exactly the edge measurements this procedure is meant to make.

- Cutouts and the square window: countersink from the back face, so the hole widens away from the sensor.

- Disks and the raised square: bevel the back, so the part narrows away from the sensor (a frustum).

- Land: the flat land left at the front edge must be ≤ EDGE_LAND_MAX_MM.

- Bevel angle: measured from the plate normal, it must exceed the worst-case ray angle plus a margin:

Chamfered cutout and disk, cross-section (not to scale)

The steepest ray past each edge travels through open space and never meets the beveled wall, so the sensor sees only the front face and the back plate.

\theta_{\mathrm{bevel}} \;\ge\; \max_{c\,\in\,\{L,\,R,\,P\}}\;\max_{e,\,\mathrm{pose}} \angle\big(\mathbf{e}-\mathbf{c},\; \hat{\mathbf{n}}_{\mathrm{plate}}\big) \;+\; \theta_{\mathrm{margin}}

Here c is the left camera, right camera, or projector center, e is any edge point at any pose used, and θ_margin = CHAMFER_MARGIN_DEG. A rough worked value, assuming about a 70° horizontal FOV and a 50 mm baseline (both unconfirmed): the worst field position at Z_MIN gives about 32° plus a 10° margin, or 42°. A standard 45° bevel (90° included countersink) would then pass. Recompute this once the VSX3000 geometry is known. If the value exceeds 45°, use a steeper bevel or reduce FIELD_OFFSET_FRACTION for the C, D, and E series.

Disk support posts. The post must sit behind the disk and be thinner than POST_DIAMETER_FRACTION_OF_D0 × the pilot D_0. For the smallest disks this means posts well under 1 mm, such as hypodermic tubing or wire. The post-only control sites confirm the post is not detected. An alternative for the smallest disks is suspension on taut sub-resolution wires, also with wire-only control sites.

As-built record. Measure each feature's front-face diameter, land width, bevel angle, and plate position with an optical comparator or calibrated microscope. Record each value with its measurement uncertainty in targets_asbuilt.csv. All analyses use these as-built values, never the nominal ones.

## Part I — Data acquisition procedure

### 4. Setup, warm-up, and robot-to-sensor registration

Registration gives every later capture a ground-truth target pose in the sensor frame. Two things anchor it. Robot repeatability fixes relative motion. An IR-image fiducial solve fixes absolute pose, and that solve does not use the stereo depth pipeline.

Experimental setup, side view

After Step 4.7, the robot-base-to-sensor transform is known, so every commanded target pose is also a ground-truth pose in sensor coordinates.

- Environment. Close the enclosure and set the lighting to its fixed state. Log ambient IR and air temperature. Keep these unchanged for the whole session.

- Sensor configuration. Disable auto-exposure and fix exposure, gain, emitter power, and trigger mode. Record every depth-processing setting (temporal filter, spatial filter, hole filling, confidence threshold) in sensor_config.json. Characterize the configuration that production will use. If the filters can be switched off, run A and B a second time with the filters off, so the sensor's own processing can be separated from its physics.

- Warm-up. Power the sensor for at least WARMUP_MIN_MINUTES. Then put T2 fronto-parallel at Z = 750 mm and capture 10 frames every minute. Start testing once the mean plane Z has drifted less than WARMUP_DRIFT_FRACTION_OF_SIGMA × σ_t over WARMUP_DRIFT_WINDOW_MIN.

- Settle and vibration check. With T2 at Z_MAX, capture 100 frames twice: once with the servos on after a move and the settle wait, and once with the brakes engaged. If the servo-on σ_t exceeds the brakes-on σ_t by more than 10%, raise ROBOT_SETTLE_TIME_S or stiffen the target mount, then repeat.

- Intrinsics and frame checks.

  - Read the left IR intrinsics, the depth-to-IR extrinsics, SENSOR_BASELINE_MM, DEPTH_LSB_MM, and the projector offset from the SDK or datasheet. Fill in the † entries in Section 2.

  - Confirm the depth image is registered to the left IR image. With T3a in view, overlay the square's edges from the IR image on the depth discontinuities. They must agree within 0.5 px.

  - Confirm the baseline direction. The occlusion band (a strip of no-reads) appears beside vertical edges only if the baseline runs along H. If it appears beside horizontal edges, swap H and V in Procedures B and E.

- Registration capture. Mount T1. Command REGISTRATION_POSES poses that span Z_MIN to Z_MAX, cover the field of view, and tilt within REGISTRATION_TILT_RANGE_DEG about H and V. At each pose capture REGISTRATION_FRAMES frames: the IR image, the depth image, and the read-back robot pose.

  - If the SDK can switch off the emitter or switch on a flood illuminator, capture the IR fiducial frames that way. The projected dot pattern degrades corner detection.

  - If neither is possible, register from depth-plane fits alone. In that case a constant depth offset cannot be told apart from registration translation (Section 15).

- Registration solve and acceptance. Solve the hand-eye problem for the camera→robot-base transform and the plate→flange transform: a closed-form start, then joint nonlinear least squares on fiducial reprojection error. Accept if the RMS distance between the solved plate poses and the measured plate poses is ≤ REGISTRATION_RESIDUAL_ACCEPT_MM. Save both transforms in registration.json.

- Station targets. For every station in Sections 5–8, compute the robot pose that puts the target's reference point at the required (H, V, Z) in the camera frame, fronto-parallel unless stated otherwise. Re-mounting a target on the dowel-pinned adapter needs no new registration. Check it once per mount by fitting the mounted target's front plane at Z = 750 mm and comparing the fit with the registered pose.

### 5. Acquisition A: noise-plate series

This series captures 11 Z stations × 5 field positions × FRAMES_PER_NOISE_STATION frames, plus a tilt sub-series. The station order is randomized so that slow drift cannot masquerade as a Z dependence.

- Mount T2. Build the station list from every Z in Z_MIN : Z_NOISE_STEP_MM : Z_MAX at five field positions, all fronto-parallel. The positions are center, then (±, ±) FIELD_OFFSET_FRACTION of the half-FOV. Check that the plate fully covers the analysis region of interest at every station. At Z_MIN off-axis this may limit the field offset.

- Shuffle the station list with a logged random seed.

- Before the first station, capture a drift sentinel: center, Z = 750 mm, 30 frames. Repeat the sentinel every DRIFT_SENTINEL_INTERVAL_MIN and after the last station.

- At each station: move, wait ROBOT_SETTLE_TIME_S, then capture FRAMES_PER_NOISE_STATION consecutive frames. Log the read-back robot pose, the sensor temperature, and timestamps.

- Tilt sub-series. At the center and Z in Z_REDUCED_STATIONS_MM, tilt T2 about V, then about H, through each angle in TILT_ANGLES_DEG. Capture 50 frames per pose. Incidence angle changes both the per-pixel noise and the fill rate.

- Repeat-mount check. Dismount T2, re-mount it, and repeat the Z = 750 mm center station. The difference shows how much of the bias comes from re-mounting the target.

- If Step 4.2 calls for it, repeat Steps 1–5 with the sensor's filters off.

### 6. Acquisition B: edge-target and Z-step series

The edge series gives lateral (H, V) resolution and most of the boundary-bias data. The Z-step series gives depth resolution. Both use robot motion: lateral sub-pixel offsets for edges, and small Z steps for depth.

6.1 Edge series (B-HV, also feeds E)

- Mount T3a (raised square) with G = GAP_SMALL_MM. Rotate it in its own plane by EDGE_SLANT_DEG, so all four edges are slanted relative to the pixel grid. The two near-vertical edges measure H resolution and the two near-horizontal edges measure V. Left and right edges have opposite occlusion geometry relative to the baseline, and so do top and bottom.

- At each Z in Z_SHAPE_STATIONS_MM, centered and fronto-parallel: move, settle, and capture FRAMES_PER_EDGE_POSE frames at the nominal pose.

- At the same Z, capture PHASE_JITTER_POSES_EDGE further poses. Each adds a logged random lateral offset, drawn uniformly over ±PHASE_JITTER_SPAN_PX/2 in both H and V. Capture FRAMES_PER_EDGE_POSE frames at each. These poses sample every sub-pixel phase of each edge for the robot-stepped ESF and for E.

- Repeat Steps 2–3 with G = GAP_LARGE_MM. Comparing the two step heights tests whether the normalized ESF depends on step height. If it does, the depth pipeline is nonlinear and resolution must be quoted together with its step height.

- Repeat Steps 1–4 with T3b (square window). The window has the opposite edge polarity: front plate outside, back plate inside.

6.2 Z-step series (B-Z)

- Mount T2 centered and fronto-parallel at each Z₀ in Z_REDUCED_STATIONS_MM. Approach every visit from below: overshoot by Z_STEP_APPROACH_OVERSHOOT_MM and move up onto the pose, so that backlash and compliance do not enter the A→B difference. The read-back robot pose, carried into the camera frame through the registration, is the step truth; log it with at least 0.01 mm resolution.

- Step ladder. For each δ in Z_STEP_LADDER_MM, alternate the target between Z₀ and Z₀ + δ for Z_STEP_REPEATS cycles (A, B, A, B, …). Capture FRAMES_PER_ZSTEP_POSE frames at each visit. Log the read-back pose at every visit; analysis uses the read-back displacement, not the commanded δ, as the step truth. Rungs smaller than ROBOT_REPEATABILITY_MM would be captured but flagged, because their truth is no better than the robot itself; the ladder therefore starts at twice that value. Alternating cancels linear drift.

- Fine staircase. Sweep from Z₀ to Z₀ + 3 × the expected quantum δZ_q (from Section 2, using the Tier-A q until A gives a measured value). Use steps of δZ_q / Z_STAIRCASE_SUBDIVISION, with 10 frames per step and the read-back pose logged at each. This reveals whether the output moves in quantized steps, or is smoothed by the sensor's interpolation.

- Blank windows. No extra capture is needed: the Z₀ frames from Step 2 serve as the no-step reference for the B-Z false-alarm threshold.

### 7. Acquisition C: disk and cutout area series

Each array is captured at many random sub-pixel offsets. Sensed area depends on where a feature's edge falls relative to the pixel grid and the projector dots. The procedure averages over that phase and also measures its spread.

- Build the configuration list: {T4-S, T4-L, T5-S, T5-L} × {GAP_SMALL_MM, GAP_LARGE_MM} × Z_SHAPE_STATIONS_MM. Center the array and keep it fronto-parallel. Randomize the order within each mounting, so targets are re-mounted as rarely as possible.

- At each configuration, capture PHASE_JITTER_POSES_AREA poses. Each has a logged random lateral offset uniform over ±PHASE_JITTER_SPAN_PX/2 in H and V, and gets FRAMES_PER_AREA_POSE frames.

- Field sub-series. At Z = 750 mm and G = GAP_SMALL_MM, repeat Step 2 for each array at the four off-axis field positions, using 10 poses each.

- Open-background variant (cutouts, optional). At Z = 750 mm, remove the back plate so that nothing lies within the sensor's range behind the holes. Then repeat Step 2. This separates the sensor's fill-in behavior from reads of a real back surface.

- Capture drift sentinels at the interval in DRIFT_SENTINEL_INTERVAL_MIN.

### 8. Acquisition D: detection trials

A detection trial is one frame at a fresh random lateral offset. All features and blank sites sit on one plate, so a single frame yields one trial for every diameter at once. The 10% and 0% minimums need two things the 50% point does not: levels that reach down to where detection disappears, and many more trials at those levels.

- Pilot. Run the quick-look detection count (Section 13, Step 2) on the first frame of each C pose at Z = 750 mm. From the counts, estimate a pilot D_50 and a pilot D_0 for each configuration. If any post-only control site shows detections above the false-alarm rate, re-make the disk posts thinner and repeat the pilot.

- Level selection. Choose DETECTION_LEVELS diameters spanning DETECTION_LEVEL_LOW_FACTOR to DETECTION_LEVEL_HIGH_FACTOR × the pilot D_50. Check that at least two levels fall below the pilot D_0. The 0% estimate needs levels with no detections, and the 10% estimate needs points on the lower tail. If the √2 ladder is too coarse to give that many levels, fabricate a dedicated detection plate with ratio 2^(1/4) across the span, built like T4 and T5 with the same blank and control sites.

- Main trials. For each configuration in {disk, cutout} × {GAP_SMALL_MM, GAP_LARGE_MM} × Z_SHAPE_STATIONS_MM, capture DETECTION_TRIALS_PER_LEVEL poses. Each pose has a new random lateral offset and FRAMES_PER_DETECTION_TRIAL frames. Shuffle the pose order with a logged seed.

- Extended trials for the 0% point. At each Z in Z_REDUCED_STATIONS_MM, increase the pose count to DETECTION_ZERO_TRIALS for every configuration. Every level, and every blank site, then has that many trials. With zero detections in n trials, the one-sided 95% upper bound on detection probability is about 3/n. So 300 trials bound a "never detected" level at ≤ 1% (DETECTION_ZERO_PROBABILITY_BOUND).

- Independence rule. Never use two frames from the same pose as separate trials. The fixed projector pattern and the target's fixed position make them strongly correlated. Section 13, Step 1 checks independence after the fact.

- Reuse. The first frame of each C pose at a matching configuration may count as a D trial, which reduces the number of new poses needed.

- Optional continuous-angle variant. Fix one diameter near the pilot D_50. Sweep Z from Z_MIN to Z_MAX in 10 mm steps, with 20 random-offset poses per step. This samples subtended angle continuously, but it mixes the angle effect with the noise change over Z; Section 13 separates them.

### 9. Data logging, file layout, and capture budget

Save every frame raw (.mc), with one manifest row per frame, so that any analysis can be re-run without re-capturing. The manifest records the read-back robot pose, not only the commanded one.

Folder layout (under Noise Estimation Data/):

Characterization_<YYYYMMDD>/
  sensor_config.json      registration.json      targets_asbuilt.csv
  manifest.csv            session_log.md         environment_log.csv
  00_registration/   A_noise/   B_edges/   B_zstep/   C_area/   D_detect/   sentinels/

File name: <proc>_<target>_G<gap>_Z<zzzz>_F<field>_P<pose>_f<frame>.mc, for example C_T5S_G15_Z0750_F0_P017_f03.mc.

Manifest columns: file name; procedure; target ID; gap G (mm); station Z (mm); field position; pose index; frame index; random seed; commanded lateral offset (H, V mm); read-back robot pose (6 values); target pose in the camera frame from registration (6 values); timestamp; sensor temperature; air temperature; ambient IR; sensor-configuration ID.

Capture budget. The estimate below assumes 10 frames/s and 3 s per move plus settle. Both are assumptions; the VSX3000 frame rate in the chosen trigger mode should replace them. Target swaps, warm-up, and the D pilot analysis are excluded. Allow two to three working days in total.

| Series | Poses | Frames | Robot time (h) |
| --- | --- | --- | --- |
| Registration | 30 | 300 | 0.03 |
| A main + tilt + sentinels | 90 | 7,030 | 0.3 |
| B-HV edges | 520 | 15,600 | 0.9 |
| B-Z ladder + staircase | 450 | 4,500 | 0.5 |
| C main + field + open | 1,420 | 14,200 | 1.6 |
| D main | 1,200 | 1,200 | 1.0 |
| D extended (0% series) | 2,880 | 2,880 | 2.5 |
| Total | 6,590 | 45,710 | 6.8 |

The extended 0% series takes about a third of the robot time. If time is short, bound the 0% point at 3/150 = 2% by setting DETECTION_ZERO_TRIALS to 150, which roughly halves that row. Check storage before starting: multiply the size of one .mc frame by about 47,000 frames.

## Part II — Data analysis procedure

### 10. Analysis A: noise vs Z

The noise is split into three parts: temporal (frame-to-frame), fixed-pattern (static, tied to the projector pattern and pixel position), and total. The temporal part is fitted to a disparity-noise model, which yields the forward-model parameters.

- Load and mask. Convert depth codes to mm with DEPTH_LSB_MM and set no-reads to NaN. Project the plate outline into the image using the registered pose. Shrink it by BOUNDARY_BAND_HALF_WIDTH_PX to form the region of interest (ROI).

- Ground truth. For each ROI pixel, intersect its ray with the registered plate plane. This gives Z_GT(u, v).

- Temporal noise. For each pixel, compute σ_t(u, v) as the standard deviation over frames (ddof = 1). Report the ROI median and 95th percentile for each station. Save the σ_t map, since its structure shows the projector-dot layout.

- Fixed-pattern noise and bias. Average the frames to get Z̄(u, v). Fit a free plane to Z̄; the standard deviation of the residual is σ_fp. The mean of (Z̄ − Z_GT) over the ROI is bias(Z). Also report the angle between the fitted plane and the registered plane.

- Total noise. For each frame, compute the residual from Z_GT. σ_tot is the RMS over the ROI and frames. Check that σ_tot² ≈ σ_t² + σ_fp² + bias²; a large mismatch points to an analysis error or to non-stationary noise.

- Fill rate. Report the fraction of valid pixels in the ROI, for each station and tilt.

- Spatial correlation. Compute the normalized 2-D autocorrelation of the temporal component (Z − Z̄), averaged over frames, and of the fixed-pattern residual map. The correlation length along H and along V is the lag where the normalized autocorrelation first falls to 1/e. Compare it with the Tier-A IDW neighborhood N and with PHASE_JITTER_SPAN_PX; the span must exceed it.

- Quantization. Histogram the raw depth codes for the central 20 × 20 px over all frames. The most common spacing between populated levels is δZ_q(Z). If δZ_q equals DEPTH_LSB_MM at every Z, the output LSB is the quantizer and disparity quantization cannot be seen here; use the B-Z staircase instead. Otherwise compute q = δZ_q · k / Z² at each Z. It should be constant if the quantizer works in disparity.

- Noise model. Fit σ_t(Z) by weighted least squares. The first term, σ_0, is a floor that does not depend on Z (LSB/√12 or electronics). The second is disparity noise σ_d propagated through d = k/Z. Also fit a free power law σ_t = a·Z^n. Triangulation predicts n ≈ 2, and a clearly different n is a finding worth reporting.

\sigma_t(Z) = \sqrt{\sigma_0^{2} + \left(\frac{\sigma_d\, Z^{2}}{k}\right)^{2}}

- Incidence angle. From the tilt sub-series, plot σ_t, σ_tot, and fill rate against incidence angle at each Z.

- Drift correction. Plot the sentinel mean Z against time and sensor temperature. If the drift exceeds WARMUP_DRIFT_FRACTION_OF_SIGMA × σ_t, subtract a time-interpolated sentinel offset from the bias estimates and report the size of the correction.

- Continuity with earlier data. Also compute the existing testZRepeatabilityBrownBoard.py metrics (20 × 20 px boxes: mean min–max range and mean per-pixel std) at Z = 700 mm and 1000 mm. This lets the new results be compared directly with the existing VSX3000 data sets.

- Outputs. A_noise_summary.csv with one row per station. Figures (PNG and SVG): the three σ curves vs Z with the model fit; fill rate vs Z; σ_t maps; autocorrelation profiles; depth-code histograms. The fitted σ_d, σ_0, q, k, and correlation length go in forward_model_parameters.json for the Tier-A simulator.

### 11. Analysis B: effective resolution in H, V, Z

Lateral resolution is read from the depth ESF across each knife edge. Two estimates are made and cross-checked: one from the slanted-edge fit, one from the robot-stepped poses. Z resolution is the step size detected with 50% probability under the same false-alarm rule used in D.

11.1 H and V resolution

- Reference levels. For each edge and frame, fit planes to the front and back surfaces. Use only pixels farther than BOUNDARY_BAND_HALF_WIDTH_PX from the edge. These give Z_front and Z_back.

- Normalized height. For each valid pixel, compute h = (Z_back − Z) / (Z_back − Z_front). Then h = 1 on the front surface and h = 0 on the back surface.

- Signed distance. Compute each pixel center's signed distance s (px) to the true edge line, projected from the registered pose. s is positive on the front-surface side.

- Robot-stepped ESF. Pool all frames and phase-jitter poses for that edge, Z, gap, and polarity. Bin by s at BOUNDARY_BIN_WIDTH_PX. ESF(s) is the mean h of the valid pixels in each bin. Keep the per-bin count of no-reads for Section 14.

- Slanted-edge ESF. On the nominal-pose mean frame, locate the half-height crossing in each row (or column) and fit a line to the crossings. Then bin by distance to that fitted line. The robot-stepped and slanted-edge ESFs should agree to within one bin once aligned. If not, suspect the registration or a curved edge.

- Metrics. The 10–90% rise distance is s(h = 0.9) − s(h = 0.1), reported in px and in mm via p(Z). The LSF is dESF/ds, smoothed with a Hann window. The MTF is |FFT(LSF)| normalized to 1 at zero frequency. Report MTF50 in cycles/px. Rise distance is the primary figure, because a depth pipeline with interpolation is not a linear shift-invariant system, and an MTF assumes one.

- Linearity test. Compare the ESFs at GAP_SMALL_MM and GAP_LARGE_MM. If they differ by more than 0.1 in h at any s, mark the result step-height dependent and quote resolution together with G.

- Edge offset. The offset between the fitted edge line and the true one, s_50 = s(h = 0.5), is the edge bias. Section 14 uses it as the near/far preference.

- Report for each Z, orientation (H or V), polarity, and gap. Confidence intervals come from bootstrapping over poses (BOOTSTRAP_RESAMPLES).

11.2 Z resolution

- Sensed step. For each consecutive A→B visit pair, compute Δ = mean(Z at B) − mean(Z at A). Do this for patches of 1 px, 5 × 5 px, and 20 × 20 px, tiled over the ROI. The true step δ is the difference of the read-back target pose along the optical axis between the two visits; the registration's translation cancels in the difference, so only the robot's relative motion accuracy enters. Rungs smaller than ROBOT_REPEATABILITY_MM are flagged as unreliable truth and left out of the gain regression of Step 2. Also report the scatter of the read-back Z over the A→A pairs as the robot's own reported repeatability.

- Gain and linearity. Regress the mean Δ on δ across the ladder. Report the slope (gain), the intercept, and the largest deviation from the line.

- Null distribution and threshold. Differences between consecutive A visits (A→A) form the no-step distribution for each patch size. Set τ at its (1 − DETECTION_FALSE_ALARM_TARGET) quantile.

- Detection curve. A step counts as detected if Δ > τ in the commanded direction. Fit the detection fraction against log δ with the psychometric model of Section 13. Report δ_50(Z, patch) with its confidence interval. For uncorrelated noise, δ_50 shrinks as the square root of patch area. Shrinking more slowly than that means the noise is correlated over that scale, consistent with the correlation length from Section 10.

- Staircase. Plot sensed Z against read-back Z for single pixels and for the 20 × 20 px patch. Plateau widths give the measured quantum δZ_q(Z). Compare it with q·Z²/k from Section 10. A smooth patch-mean curve over a stepped single-pixel curve shows that the sensor's interpolation dithers the quantizer.

### 12. Analysis C: true vs sensed area

Sensed area is compared with two references. One is the true area of the as-built feature. The other is the area that stereo geometry allows the sensor to see. The second matters for cutouts: part of the back plate seen through a hole is visible to only one camera, or is not lit by the projector, so no stereo sensor could read it.

- Feature window. For each feature, frame, and pose, project the as-built outline into the image using the registered pose. The window is a circle of radius D/2 + BOUNDARY_BAND_HALF_WIDTH_PX (in px at that Z) around the projected center.

- Reference planes. Fit the front plane to an annulus of front surface outside the window. For a cutout this is the front plate. For a disk it is the registered disk plane, because small disk faces have too few pixels to fit. Fit the back plane to an annulus of back plate beyond the shadow band.

- Pixel classes. Compute h as in 11.1. Inside the window, each pixel is a front read (h ≥ 0.5), a back read (h < 0.5), or a no-read.

- Sensed area per frame. For a disk, sensed pixels are the front reads connected to the projected center. For a cutout, they are the back reads connected to it. Each pixel's area on the feature plane is a = Z_plane(u, v)² / (f_x f_y cos φ), where φ is the incidence angle. Report three quantities: A_sensed = Σa over sensed pixels; A_noread = Σa over no-reads inside the projected outline grown by 1 px; and A_sensed + A_noread, an upper bound for a reader that treats no-reads as part of the feature.

- Sub-pixel area. For each pose, average its frames and interpolate h over valid pixels. Extract the h = 0.5 iso-contour with marching squares, deproject it onto the feature plane, and take the polygon area. Use this to check the per-frame pixel-count result for systematic error.

- True area. A_true = πD²/4, using the as-built D.

- Geometric visible area. Ray-cast from the registered left camera, right camera, and projector centers through the as-built outline to the back plane. A_geo for a cutout is the overlap of the three projected outlines. Compute it with all three, and also with only the two cameras, because it is not known whether the VSX3000 needs projector light on a point to read it. Report which version tracks the data better. For a disk front face, A_geo = A_true.

- Transfer curves. Plot A_sensed / A_true and A_sensed / A_geo against D_px = D·f_x / Z, for each Z, gap, and feature type. Put disks and cutouts on the same axes so their asymmetry shows. Report the mean, the spread across poses (phase effect), and the spread within a pose (temporal effect).

- Edge bias. Convert to an equivalent diameter D_s = 2√(A_sensed/π). The edge bias is b = (D_s − D)/2, in mm and in px. Fit b against Z using diameters well above D_50, where b should not depend on D.

- Consistency with B. Build a 2-D blur kernel from the H and V line spread functions (assumed separable). Blur an ideal disk of the as-built D with it and threshold at 0.5; this predicts the area. Agreement for large D and disagreement for small D means the sensor's interpolation fills in small features beyond what the edge spread function explains. Also check that the sign of b matches the edge offset s_50 from 11.1.

- Field and open-background variants. Compare the off-axis transfer curves with on-axis ones. For the open-background cutouts, compare the no-read area with A_true. Fill-in by front-plate values shows as A_noread < A_true.

- Outputs. C_area_summary.csv, plus figures of the transfer curves, b vs Z, and phase-spread vs D_px.

### 13. Analysis D: minimum detectable size at 50%, 10%, and 0%

D_50 and D_10 come from a psychometric curve fitted to the detection trials, after correcting for false alarms. D_0 needs separate treatment. A smooth curve never reaches zero, and no finite number of trials can prove a probability is zero. So D_0 is reported as the largest size whose detection probability is shown, with 95% confidence, to be below DETECTION_ZERO_PROBABILITY_BOUND. It is cross-checked against a model with a hard floor and, for cutouts, against the geometric limit.

- Independence check. For each feature, compute the lag-1 autocorrelation of its detected/missed sequence in acquisition order. It should lie within ±2/√n. Also check that outcomes at neighboring features in the same frame are uncorrelated. Correlation at either level means the offsets were too small or the features too close; widen PHASE_JITTER_SPAN_PX or FEATURE_ISOLATION_PX and repeat.

- Detection rule. In the feature window (radius D_px/2 + DETECTION_WINDOW_MARGIN_PX), compute the deviation from the reference plane. For a disk, Δ = Z_back − Z (toward the sensor). For a cutout, Δ = Z − Z_front (away from the sensor). Candidates are valid reads with Δ > τ. The feature counts as detected if at least DETECTION_MIN_CONNECTED_PX connected candidates exist. No-reads do not count in the primary rule. A second, "no-read-inclusive" rule treats no-reads as candidates for cutouts, since a downstream reader might take a hole of no-reads to be a hole. Report both, because the gap between them is a direct measure of boundary bias (Section 14).

- Threshold. For each Z, configuration, and window size, apply the rule to the blank sites. Set τ so that the false-alarm fraction equals DETECTION_FALSE_ALARM_TARGET. Then record the measured false-alarm rate γ with its Clopper–Pearson interval.

- Psychometric fit. Model the raw detection probability as below, with x = ln D. F is the logistic function, γ is fixed from the blank sites, and the lapse rate λ is free but bounded to [0, 0.05]. Fit α, β, λ by maximum binomial likelihood. The false-alarm-corrected probability is P*(D) = F.

\psi(D) = \gamma + (1-\gamma-\lambda)\,F\!\left(\frac{\ln D - \alpha}{\beta}\right), \qquad D_{50} = e^{\alpha}, \qquad D_{10} = e^{\alpha - \beta \ln 9}

- Model dependence of D_10. D_10 sits in the tail, so it depends on the curve shape more than D_50 does. Fit three shapes for F: logistic, cumulative normal, and Weibull in ln D. Report D_10 from the best fit by deviance, and the range across the three as model uncertainty. As a check that needs no model, fit a monotone curve to the corrected proportions with isotonic regression and interpolate the 0.1 and 0.5 crossings.

- D_0, empirical bound (primary). At each level, compute the one-sided Clopper–Pearson upper bound ψ_U on the raw detection rate at CONFIDENCE_LEVEL. Correct it: P*_U = (ψ_U − γ)/(1 − γ). D_0,emp is the largest level for which P*_U ≤ DETECTION_ZERO_PROBABILITY_BOUND holds at that level and at every smaller level. Report the pair [D_0,emp, next level up]. The true 0% point lies between them, so its resolution is the ladder spacing. This is why Step 8.2 calls for a finer detection plate.

- D_0, threshold model (cross-check). Fit a Weibull with a location parameter: P* = 0 for D ≤ D_0, and 1 − exp(−((D − D_0)/s)^k) above it. Give D_0 a profile-likelihood confidence interval. This is an extrapolation that depends on the assumed shape. Treat agreement with Step 6 as support, and disagreement as a sign that the floor is not sharp.

- D_0, geometric limit (cutouts). Find the diameter at which A_geo from Section 12, Step 7 reaches zero, meaning no part of the back plate is seen by all the views a read needs. Below it, the primary rule cannot detect a cutout on any sensor of this baseline. A D_0,emp well above the geometric limit means the sensor's processing, not geometry, sets the floor.

- Confidence intervals. Bootstrap over poses (BOOTSTRAP_RESAMPLES), re-fitting γ, τ, and the curve each time. This gives intervals for D_50, D_10, and the model D_0.

- Units and Z dependence. Report each minimum in three forms: D in mm, D_px = D·f_x/Z, and θ = D/Z in mrad. Plot θ_50, θ_10, and θ_0 against Z. A constant D_px suggests the minimum is set by pixel or projector-dot sampling. A D_px that grows with Z in step with σ_tot(Z)/G suggests it is set by noise. For the optional continuous-angle variant, fit a logistic regression on ln θ and ln σ_tot(Z) together to separate the two.

- Outputs. D_detect_summary.csv with one row per configuration and Z. Columns: γ, τ, D_50, D_10, the D_0,emp bracket, model D_0, geometric D_0 (cutouts), and confidence intervals, each in mm, px, and mrad, for both detection rules. Figures: psychometric curves with binomial error bars and D_0, D_10, D_50 marked; θ vs Z for each minimum.

### 14. Analysis E: boundary detection bias

At a boundary, the sensor can err in two directions. It can report a value where geometry says no stereo reading is possible (a fabricated read, usually interpolated). Or it can report a no-read where a valid surface was visible (a dropout). Comparing the two rates shows whether the VSX3000 leans toward reading values or toward no-reads. A second measure shows which surface it favors when it does read: the near (front) surface or the far (back) one. Both come from the B edge data and the C array data, so no new captures are needed.

- Geometric visibility mask. For every pixel in a boundary band (|s| ≤ BOUNDARY_BAND_HALF_WIDTH_PX), ray-cast its true surface point from the registered pose. Test whether that point is visible to the left camera, the right camera, and the projector. Set V = 1 if a stereo read is geometrically possible and V = 0 if not. As in Section 12, Step 7, compute V both with and without the projector condition.

- Outcome classes. Classify each pixel-frame in the band as one of four outcomes. Front read: valid and within SURFACE_ASSIGNMENT_SIGMA_MULTIPLE × σ_tot(Z) of the front plane. Back read: the same test against the back plane. Intermediate read: valid but near neither plane (a "flying pixel"). No-read.

- Outcome profiles. Bin by signed distance s to the true edge, at BOUNDARY_BIN_WIDTH_PX, with s > 0 on the front side. Plot P_front(s), P_back(s), P_mid(s), and P_none(s), which sum to 1. Use radial distance r − D/2 for disks and cutouts. Pool over phase-jitter poses so every sub-pixel phase is equally represented.

- Read vs no-read bias. Compute the fabrication rate R_fab and the dropout rate R_drop as below. Combine them into a bias index β_read from −1 (always drops out) to +1 (always fills in). Also report two widths in px: W_fab, how far reads extend into the geometrically unreadable region, and W_drop, how far no-reads intrude into the readable region. Widths are easier to use in a forward model than rates.

R_{\mathrm{fab}} = P(\mathrm{read} \mid V=0), \quad R_{\mathrm{drop}} = P(\mathrm{no\text{-}read} \mid V=1), \quad \beta_{\mathrm{read}} = \frac{R_{\mathrm{fab}} - R_{\mathrm{drop}}}{R_{\mathrm{fab}} + R_{\mathrm{drop}}}

W_{\mathrm{fab}} = \int_{V=0} P_{\mathrm{read}}(s)\, ds, \qquad W_{\mathrm{drop}} = \int_{V=1,\;|s| \le s_{\mathrm{band}}} P_{\mathrm{none}}(s)\, ds

- Near vs far preference. Count wrong-surface reads: front reads at s < 0 (N_fb) and back reads at s > 0 (N_bf). The index π_near = (N_fb − N_bf)/(N_fb + N_bf) runs from −1 (favors far) to +1 (favors near). Stereo matchers with a correlation window usually show positive π_near ("foreground fattening"), but the VSX3000's behavior is not known in advance. Cross-check against the edge offset s_50 from 11.1, Step 8, and the sign of b from Section 12. Also report the mean h of intermediate reads: 0.5 means even blending, and values near 0 or 1 mean blending that leans to one surface.

- Feature-scale bias. For each cutout diameter, classify the region inside the true outline by majority outcome in each frame: back read (correct), front read (fill-in), or no-read. Plot P(fill-in), P(no-read), and P(correct) against D_px. The fill-in fraction near D_50 shows directly whether the sensor reports a solid surface or a hole where it cannot resolve. For disks, the matching measure is the fraction of disks reported as back plate (erased) versus no-read. Also report the difference between the two detection rules of Section 13, Step 2, at each D.

- Breakdowns. Report β_read, W_fab, W_drop, π_near, and s_50 for each edge orientation (along the baseline vs across it), polarity, Z, and G, and for straight edges vs circular features. Confidence intervals come from bootstrapping over poses.

- Outputs. E_boundary_bias.csv, plus figures: stacked outcome profiles vs s for each edge type; β_read and π_near vs Z; feature-scale outcome fractions vs D_px. The widths and π_near go into forward_model_parameters.json as boundary terms for the forward model.

### 15. Uncertainty budget, limitations, and open questions

The smallest Z steps and the absolute bias are the measurements most limited by the setup. The relative measurements are robust: noise, rise distance, transfer-curve shape, and detection minimums. The magnitudes below are typical values, not measured ones; replace them with measured values after Step 4.

| Source | Estimated magnitude | Affects | Mitigation |
| --- | --- | --- | --- |
| Robot repeatability | 0.02–0.05 mm | Smallest B-Z steps; edge position | Approach every Z visit from the same direction; A, B, A, B averaging over the cycles; rungs below the robot repeatability flagged, and δ_50 reported as a bound when it falls below them; averaging over phase poses |
| Robot absolute accuracy | 0.2–1 mm over large moves | bias(Z), registration | IR fiducial solve over 30 poses; tests run as local moves from registered stations |
| Registration residual | ≤ 0.15 mm RMS (acceptance limit) | bias(Z), edge offset s_50, small-feature area | Acceptance gate (Step 4.7); report the residual with every result |
| IR fiducial depth accuracy at 1 m | a few tenths of a mm (rough estimate) | Absolute bias(Z) only | Many poses; plate pattern certified to ≤ 0.02 mm |
| As-built diameter uncertainty | set by the measuring instrument | A_true: relative error ≈ 2δD/D, for example 8% at 0.5 mm with δD = 0.02 mm | Optical comparator or microscope; carry δD into the area and D_0 confidence intervals |
| Plate flatness | ≤ 0.05 mm | Fixed-pattern noise, bias | Flatness spec; map the plate on a CMM if available |
| Thermal drift | unknown until the sentinels run | bias(Z), B-Z | Warm-up gate, sentinels, ABAB ordering, randomized station order |
| Front–back interreflection | unknown | Cutout and small-gap results | Compare G small vs large; matte finish |
| Sensor temporal filter | depends on configuration | σ_t (underestimated), trial independence | Discard frames after each move until the plane mean settles (set in Step 4.4); run filters off if possible |
| Reflectance mismatch | avoided by design | Edges, area | One finish on all surfaces |

Limitations. Results hold for one surface finish, static targets, mostly fronto-parallel poses, and the sensor configuration recorded in sensor_config.json. If the IR fiducial solve is not possible (Step 4.6.2), a constant Z offset is absorbed into the registration and absolute bias(Z) is not measurable. Z-scale and nonlinear bias still are, because a rigid transform cannot absorb them. A "0%" minimum means below the stated bound with 95% confidence, not proven zero. Because the read-back robot pose is the step truth of B-Z, the smallest detectable Z step for the larger patches may come out as a bound set by the robot's repeatability rather than a measurement.

Open questions

- ☐ VSX3000 f_x, f_y, baseline, projector offset, FOV, depth LSB, and frame rate in the chosen trigger mode. These fill the † parameters, the chamfer check, and the time budget.

- ☐ Can the emitter be switched off, or a flood illuminator switched on, for fiducial imaging? This decides between the independent and the depth-only registration.

- ☐ Which on-sensor processing configuration is the production one?

- ☐ Which surface finish represents the production targets, and is a second finish needed?

- ☐ Does the robot controller report the actual pose at the time of capture?

- ☐ Does the robot write its read-back pose with at least 0.01 mm resolution, and is it the actual (encoder) pose rather than the commanded one? For B-Z the read-back pose is the step truth, so a commanded-only report makes the rungs below the robot's accuracy meaningless.